Blood Gas Analyser

24/11/10

USES – point of care measurement of arterial blood gas indices + electrolytes

DESCRIPTION

Direct measurements

- pH

- PCO2

- PO2

Derived measurements

- HCO3-

- BE

- O2 saturation

METHOD OF INSERTION AND/OR USE

pH measurement

- glass electrode + reference electrode

- specimen put in a capillary tube surrounded by buffer solution

- the tube is made of pH sensitive glass across which a potential difference is generated -> this is proportional to the pH

PCO2 measurement

- modified glass electrode

- comprises of a glass pH electrode that is permeable to CO2

- CO2 diffuses from the specimen into the HCO3- solution where it dissociates with a change in pH which is measured by the electrode.

- potential difference is proportional to CO2 concentration

PO2 measurement

- Clark electrode or polargraphic electrode

- O2 molecules diffuse across a plastic membrane -> small platinum or gold 2nm diameter wire cathode in a glass rod immersed in a phosphate buffer with KCl

- O2 reduced by 2 hydroxyl ions by 4 electrons after application of 600-800mV

O2 + 2H2O + 4e = 4(OH)

- a resultant current flow between the cathode & silver/silver chloride anode.

- increasing the voltage increases the current up to a plateau determined by the rate of supply of O2 molecules & proportional to the concentration of O2 in solution.

OTHER INFORMATION

HCO3-

- derived from transposition of Hendersen-Hasselbalch equation

[HCO3-] = [K [0.03] x PaCO2] / [H+]

where K is the combined CO2 solubility constant and carbonic acid dissociation constant which can vary with pH, temp & ionic strength of solution.

- calculated actual HCO3- is assumed to by 24mmol/L

BE

- titratable acid in mmol/L required to titrate blood to a pH of 7.4 at a PCO2 of 40mmHg

- used to assess what changes in pH are due to PaCO2 changes.

- ie. at a pH of 7.4 the change in [HCO3-] from normal bicarb = the amount of HCl added.

- normal +2.5 to -2.5

- large negative values = metabolic acidosis

O2 saturation

- PaO2 ratio of oxyHb to total Hb

- normally 95-98%

COMPLICATIONS